Digital Storage Oscilloscope and Logic Analyzer

The device described in this article is a combination Digital Storage Oscilloscope and Logic Analyzer which interfaces to a PC through a parallel port.  It is built using several printed circuit boards.  The system is designed such that it allows for 6 Input Cards.  Each card can be either 8 Logic Analyzer channels or a single Analog channel (8 bit resolution).

A previous article ( QEX, Jan 2000 ) describes an 8 channel Logic Analyzer which I built and have used quite a bit.  The knowledge and experience gained in its development were instrumental in the development of this project.  The original intent was for a DSO only; however, it was soon obvious that most of the design was also useful for implementing a Logic Analyzer.  I used many of the ideas from the LA in the design of this combination device.

My main goals for the DSO/LA were:

(1) ease of construction - use printed circuit boards

(2) two analog channels

(3) minimum of 20 MHz sample rate

(4) 8 Logic Analyzer channels

(5) trigger selection to allow any combination of analog and LA signals

The current design consists of printed circuit boards which stack one upon another.  They are connected together via “board stacking” connectors.  This method allowed me to design a system which is easy to modify, expand and customize.  It allows the builder to implement up to 48 Logic Analyzer (in groups of 8) or a maximum of 6 Analog channels and any combination in between.  The resulting assembly is fairly close to a 4.5 inch cube with two Analog cards and one Logic Analyzer card.  Great care was taken in the design of the printed circuits in order to achieve this modularity.  A template board was first designed, having just the connectors, which was then used as the basis for all the boards.  The relative stacking positions of the boards was chosen to make maximum use of the connectors for inter-board connections.

I wanted to achieve a 20 MHz sample rate so that the analog channels would be useful to 10 MHz.  The A/D converter selected (ADS820) is specified to a 20 MHz sample rate, 60 MHz bandwidth, and has 10 bit resolution.  Since the circuit uses only the most significant 8 bits some of the errors of the A/D can be ignored.  There is a higher frequency A/D available which is pin compatible with the unit I selected but I felt that I would not be able to implement a circuit board which would allow the additional bandwidth.

Major Circuits

Control - Interface to the Printer Port, source for all the static control signals

Trigger Address - stores the address at which the trigger occurs

Trigger Control - Develops most of the dynamic control signals, Pre/Post trigger count

Trigger Detect - Detects the trigger condition

Address Generator - 17 bit address = 132K

Time Base - generates the seven sample clock frequencies

Logic Analyzer - CMOS levels, 8 channels, 132K memory depth

A/D Converter - 0 -> +4 volts, 8 bit resolution using a 10 bit A/D, 132K memory depth

Circuit Descriptions

General Comments

I have followed a naming convention for the signal names in order to more easily determine their operation:

(1) dynamic signals will have a PL or PH suffix indicating Pulse Low or Pulse High

example: SysReset - PL

(2) static signals will have an H or L suffix indicating a High true or Low true state

example: Read-Ch0-L

(3) static signals which are neither high or low true do not have a suffix

example: TimeBaseSel0

(4) The ICs will have a two digit number with the first being the board/circuit number.

In general I have labeled every signal which goes between circuits.  I have not generally labeled signals which do not go between circuits.

Control Circuit

The main purposes of the Control Circuit are to provide the interface to the PC (parallel printer port) and to develop the static signals used to control the remaining circuits.  The printer port control bits are used to select the primary decoder via U11.  The functions performed by the primary decoder are:

- System Reset

- System Enable

- Secondary Decoder Latch and Time Base Select - U12

- Read Clock and Secondary Decoder Read Strobe - U15

- Secondary Decoder Write Strobe - U14

Most of the signals generated by this circuit are low true pulses.

There is also a bus transceiver (U13) which determines the direction of data flow between the PC and the DSO/LA.  It uses the ReadClock to determine the data driection: low = read from DSO/LA,  high = write to DSO/LA.

Although many of the circuits allow for customization this one does not since the software is written to correspond with the functions it performs.

Time Base

The operation of the Time Base is rather straight forward.  There are three (or four) control bits which come from the Control Circuit.  These bits control a multiplexer which is used to select from among eight signals which can clock the Address Generator (AddrClk-PH).  Seven of these are also the Sample Clock signal for both the A/D and Logic Analyzer Circuits (SampleClk-PH).  The eighth signal is for advancing the Address Generator when the program needs to read the data from the A/D and Logic Analyzer RAMs.  The 20 Mhz oscillator and Sample Clock are enabled by a signal from the Trigger Control Circuit (SampleClkEna-L).

Notice that the signal which drives the Address Generator goes through three gates, wired as simple inventors, while the signal which drives the Sample Clock only goes through a single gate.  There are two reasons for this configuration:

(1) the Sample Clock needs to be disabled when reading from the A/D and LA Circuits.

(2) The Address Clock is delayed slightly so that the input signals are sampled a short time (two gate delays) before the address bus is incremented.  This allows the input signals to settle for almost a full clock period and is really only significant at the highest sample rate but is necessary to meet the access time specification of the RAM.

I elected to make the sample rate/period circuit as simple as possible.  I have used binary dividers to generate the various sample rates.  Starting with a 20 MHz clock the next rate is 10 MHz, followed by 5 MHz, etc.  The sample periods with this circuit are: 50 ns, 100 ns, 200 ns, 400 ns, 800 ns, 1600 ns and 3200 ns.  These may seem like “strange” values but it really does not make any practical difference in the interpretation of the displayed signals.  Since the software allows you to measure the time between cursors it accounts for the sample period.  Also, the software allows you to set the time scale.

The design of the system is such that if you do not like the sample periods I have chosen you can redesign the circuit to develop the ones you like.  You can specify the sample periods in the INI file.  All you need to do is relate the 3 or 4 bit decoder input value to the sample period as follows: SELECT_<decode value:1..15> = <sample period>

For example: SELECT_1 = 50ns

The program will correctly interpret ns, us and ms as units of time.

Address Generator

This is 17 bit synchronous binary counter which yields 2^17 = 131072 addresses.  It is driven by the Address Clock from the Time Base Circuit.  The counter outputs are all set to zero by the SysReset signal from the Control Circuit.  Due to the way the Pre and Post Trigger circuits operate, the full count is not utilized.  This will be explained further in the discussion of the Trigger Control circuit.

The original design of the circuit called for 74HC161 binary counters which are specified to 25 MHz with a 5 volt supply.  However, when more that two units are cascaded there is a “glitch.”  I had not thoroughly read the details of the specification and missed the information on the glitch.  I thought that since they were rated to 25 MHz I would not have a problem since I was only going up to 20 MHz.  The glitch occurs due to internal delays when using the Ripple Carry Out.  This limits the upper frequency to 17 MHz when cascading.  The solution was relatively simple: use 74AC161 devices instead.  They are rated at a much higher frequency but draw quite a bit more current.

Trigger Control

Trigger Detect

A/D

Logic Analyzer

Trigger Address

A/D Buffer

Some general comments and suggestions for making printed circuit boards:

(1) Make the traces as wide as you can.  Most of the boards I developed started out with 20 mil traces.  After completing the design I would then go back and make each trace as wide as possible.  Where necessary I would “neck down” the trace in order to fit it between IC or connector pins.  The main reason for doing this is to prevent undercutting of the traces during etching.

(2) Make the power and ground traces even wider.  This is done to give the least resistance for the power leads.  One of the guidelines I followed was to implement as many of the signal traces as possible ignoring the power.  Since this was my first effort at double sided printed circuit boards and since I do not have the ability to implement feedthrus I felt that running wire jumpers for power and ground was better than running jumpers for signals.  This allowed me to use relatively heavy wires for the power.

(3) Implement ground “fill” where possible.  This reduces the amount of copper which needs to be etched as well is giving more ground area.

(4) Use a piece of scrap perforated board as a drilling template for the connector holes.  Lining up the drill bit by eye does not work well when you need to insert a 35 pin connector!  I had some unused circuit boards from Radio Shack which have 0.1” spaced holes.  I lined up the holes with the positions to be drilled and then taped the perf board to my circuit board.  Looking through the board with a 50 watt lamp behind it made it relative easy.

Some general guidelines I implemented:

(1) Drill as few holes as possible.  Looking closely at the photos you may notice that the IC sockets look a little odd.  I designed the boards as if I were using SMD ICs.  However, I did use mostly normal DIP devices.  I found IC sockets which have the flat side of their pins parallel to the long edge.  This allowed me to bend them to the side at a right angle to the socket.  I laid out the IC traces to insure at least 0.1” of trace so that the socket pins could be soldered in place.  I had to trim the leads on a few sockets but this is quite easy.

(2) Allow ICs on both sides of the boards where necessary.  This was easily implemented due mainly to (1) above.

(3) Keep the main circuit functions on separate boards.  This makes for a very modular system which can be easily customized.  It also makes testing the individual circuits easier.

(4) Install “load” resistors on selected inputs.  This, again, makes it easier to test the boards.

Printer Port Signal Usage

Output Control Bits

Pin	Name			Usage

1	/C0: Strobe		RegSel 0

14	/C1: Auto-Linefeed	RegSel 1

16	 C2: Initialize		RegSel 2

17	/C3: Select-Printer	Strobe

Input Control Bits

Pin	Name			Usage

10	S6: Ack			PreTrigDone-L

11	/S7: Busy

12	S5: Paper-Out		AddrClkEna-L

13	S4: Select		Trigger Latch-L

15	S3: Error

2 - 9	D0 - D7			Data

18	Ground

Bill of Materials by Circuit

Consolidated BOM