Biocompatibility considerations – Interactive Review for Emergency Use Authorization (EUA)

This document lists the biocompatibility recommendations for medical devices that are not marketed in the US or outside of US. It does not discuss the devices that are marketed for a different indication, contain a different material, contact duration, etc. and the sponsor had modified the device without the submission of a 510 (k) (refer to the EUA Ventilator March 2020 guidance). In addition, it does not discuss devices that claim conformance to applicable FDA recognized standards (refer to Appendix A of the EUA Letter)

- Regulatory history

If the device has been previously submitted for review to the FDA, provide the submission number.

- Information on device materials with direct and indirect patient contact

Please provide available information, as listed below, as well as any other relevant information that may not be covered by the items below:

- List of materials with direct contact with the patient and with the gas pathway.

- The material manufacturer has a Master file (MAF) available with the FDA.

- The material is medical grade material such as material that was evaluated per United States Pharmacopeia (USP)? Provide the USP classification (e.g., Class VI)?

- Safety Data Sheet (SDS) and/or certification of analysis (COA).

- Information available from the vendor including chemical registry number (CAS#), chemical composition and percentage, presence of impurities, mixture(s), presence of color additives, etc.? Discuss if the polymers contain the following additives, impurities, residuals, and degradation products:

    - Plasticizers, flame retardants, stabilizers, antioxidants, UV stabilizers, heat stabilizers, slip agents, lubricants, anti-statics, curing and blowing agents, biocides, colorants, fillers, fragrances, solvents, etc.

    - Unreacted monomer/oligomer

    - Degradation products

  - Statement from material supplier(s) that a material does not contain specific substances (latex, heavy metals, phthalates, polycyclic aromatic hydrocarbons, aldehydes, peroxide, animal tissue or blood components, radioactive material, etc.).

- The material manufacturer claim conformance to a material standard (e.g., American Society for Testing and Materials (ASTM))?

- Statement from the material supplier that the material is used for other US-marketed devices. If information is available, specify if the material is the same, from the same supplier, as material used in another marketed device with same category and contact duration?

- Information available (e.g., from literature, material supplier) that the material is non-reactive and stable during/after proposed sterilization method.

- Manufacturing information

- Available information on the use of any manufacturing aid (e.g. solvent, oil, lubricant, cleaning reagent, etc.), and whether any assessment was performed to determine that these are not present in the device/have the potential to reach the gas pathway and/or the patient.

- Hazard identification and toxicity assessment

- Identify hazard information from SDS, is available. For example, Globally Harmonized System Hazard Class Pictograms:

Source: A Guide to The Globally Harmonized Systemic of Classification and Labeling of Chemicals available at https://www.osha.gov/dsg/hazcom/ghsguideoct05.pdf

- Identify potential hazardous volatile organic compounds (VOC) or leachable substances that may be present in the finished device.

Information from material suppliers, standard material processing methods, and common material impurities/contaminants may be considered. Refer to the following documents for lists of hazardous air contaminants: Clean Air Act Amendments of 1990 List of Hazardous Air Pollutants (https://www3.epa.gov/airtoxics/orig189.html) and the OSHA Standard 1910.1000 Table Z-1 and Z-2 (https://www.osha.gov/laws-regs/regulations/standardnumber/1910/1910.1000TABLEZ1 and

https://www.osha.gov/pls/oshaweb/owadisp.show_document?p_id=9993&p_table=STANDARDS).

- Discuss any assessments performed to ensure that the device does not release noxious gases during use (e.g., ozone, carbon monoxide, fumes)?

- Discuss if the materials with direct contact with the patient and with the gas pathway, the manufacturing materials and any chemicals used, are known to be classified in any of the categories listed below. If material(s) had been classified in any of these categories, discuss the risk mitigations.

- Respiratory sensitizers (please refer to the list below)

additional classification information available at (if Internet Explorer does not open the links, use Google Chrome or Firefox): https://www.schc.org/assets/docs/ghs_info_sheets/respiratory_sensitization_2017-04.pdf

- Acute toxicity (please refer to the list below)?

additional classification information available at: https://www.chemsafetypro.com/Topics/GHS/GHS_classification_criteria_acute_toxicity_category.html

- Specific target organ toxicity after single exposure (please refer to the list below)?

additional classification information available at: https://www.schc.org/assets/docs/ghs_info_sheets/specific_target_organ_toxicity-single_exposure.pdf

- Serious eye damage and irritation (please refer to the list below)?

additional classification information available at: https://www.schc.org/assets/docs/ghs_info_sheets/Eye%20Damage%20Irritation%20Final%202018-03.pdf

- Skin corrosion and irritation (please refer to the list below)?

additional classification information available at: https://www.schc.org/assets/docs/ghs_info_sheets/Skin%20Corrosion%20%20Irritation%20(Final%202018-03).pdf

- Cancer (please refer to the list below)?

additional classification information available at:

https://www.schc.org/assets/docs/ghs_info_sheets/schc_osha_carcinogenicity_info_sheet_february_2014.pdf

- Skin sensitization (redness, itching skin, bumpy rash) (please refer to the list below)?

additional classification information available at:

https://www.schc.org/assets/docs/ghs_info_sheets/skin_senzitization_2017-04.pdf