Device specifications:

| Parameter | Range | Comments |
| --- | --- | --- |
| RR (Hz) | 10-40 |  |
| TV (mL) | 200-800cc | Validated only for these TV |
| Flow Rate (L/min) | 6-15L/m | Please see details on specs of validated reservoir bag. |
| PIP (cm H20) | 30-60 | Range of validated PIP |
| PEEP (cm H20) | Up to 25 | Validated PEEP range |
| Available Ventilation Modes | PCV, VCV, PSIMV, VSIMV, PS |  |
| Air Source | Internal blower | Technical specifications |

Labeling Parameters:

| Parameter | ARDSnet Conformance | Comment |
| --- | --- | --- |
| PC/VC | VC only | Can only be operated in VC mode |
| 4-6 cc/kg TV | Weight <115.9 kg | *Calculated for 6cc/kg limit |
| FiO2 (.21-1.0) | No* | *Please note this is an estimate and calculation by flow rates |
| I:E ratio (1:1-1:3) | Yes* | *Can do 1:4 |
| RR (up to 35) | Yes | RR max 40 |
| PEEP (up to 24 cm H2O) | Yes |  |
| Pplateau | No | ARDSnet goal Pplateau ≤ 30 cmH2O |
| Validation and Performance Conformance |  |  |
| Compatible accessories | Humidifier, breathing circuits | Met ISO 80601-2-74 |
| Accuracy of Monitored parameters (as applicable) | Volumes ± 10% | Characterized per ISO 80601-2-80:2018 |
|  | Pressures (2 + 4 % of the actual reading) hPa (cmH2O) | Met ISO 80601-2-80:2018 |
| Accuracy Delivered (as applicable) | Volumes ± 20% | Characterized per ISO 80601-2-80:2018 |
|  | Pressures ± 10% | Characterized per ISO 80601-2-80:2018 |
| Spontaneous Breathing detection | Yes | Flow trigger |
| Internal power source ventilation duration capacity | 30 minutes | Back-up Battery checked at maximal device performance for 30 minutes |
| Contamination risk to environment | No* | *Need HMEF (99.97% efficiency at ETT-Wye adaptor interface. |
| Durability | Yes | tested to 1 week |
| Alarms |  |  |
| Power-off Alarm | Yes | Continuous audible alarm needing manual disable upon powering-off |
| Disconnect Alarm (Low PIP alarm) | Yes (within two breaths)* | LED and audible alarm, per ISO 80601-2-80:2018 *within two breaths |
| External power supply failure Alarm | Yes (10 seconds)* | LED and continuous audible alarm that needs manual disable by clinician, per ISO 80601-2-80:2018 *within 10 seconds |
| Back-up Battery | Yes* | Automatic, and triggered by power supply failure alarm at 10 sec. , per ISO 80601-2-80:2018 *continuous audible Alarm that needs manual disable by clinicians 5’ prior to battery depletion. |
| Low FiO2 alarm | No | No FiO2 alarm. Requires vigilance to ensure adequate oxygen flow rates to specify needed FiO2 |
| Maximum PIP alarm | Yes at 40cm H2O | LED and audible alarm @40cm H2O, per ISO 80601-2-80:2018 |
| Occlusion alarm/Continuing Pressure | Yes | LED and audible alarm, per ISO 80601-2-12:2020 |
| TV not met/ exceeded | Yes | LED and audible alarm, per ISO 80601-2-12:2020 |
| Hypoventilation | Yes | LED and audible alarm, per ISO 80601-2-80:2018 |
| PEEP Alarm | Yes | LED and audible alarm, per ISO 80601-2-12:2020 |
| Leakage Alarm | Yes | LED and audible alarm, per ISO 80601-2-80:2018 |
| Implications for additional clinical monitoring | Standard ARDS clinical monitoring + continuous end-tidal CO2 monitoring* | *continuous end-tidal monitoring indicated since no alarm for TV not met/exceeded |
| Training |  |  |
| Training required | 1 hour of minimal training |  |

Acronyms: PC – Pressure Control, VC – Volume Control, FiO2 – Fraction of Inspired Oxygen, I:E ratio – inspiration – expiration ratio, RR – respiratory rate, PEEP – positive end-expiatory pressure, Pplateau – Plateau Pressure, PIP – peak inspiratory pressure, TV -tidal volume